Active Birth Pools

Cleaning & Disinfection Protocol

For Ficore® Birth Pools in US and Canadian Healthcare Facilities

Recommendation for Infection Control Personnel

# 1.  Pool Construction & Material Compatibility

Active Birth Pools are fabricated from three materials, each with distinct chemical compatibility profiles that must be considered when selecting cleaning and disinfection agents:

| Material | Description | Chemical Sensitivity |
| --- | --- | --- |
| Ficore® composite | Proprietary isophthalic neo-pentyl-glycol composite; primary pool basin surface with which patients and water contact occurs. | Sensitive to abrasives and concentrated oxidizing acids; compatible with most neutral- to mildly-alkaline disinfectants at recommended dilutions. |
| Stainless Steel | Grab handles, hardware. | Susceptible to chloride-induced pitting and stress corrosion cracking, particularly with prolonged or concentrated chlorine-based products. Performs well with quaternary ammonium and accelerated hydrogen peroxide products. |
| Chrome-Plated Brass | Drain assembly. | Most chemically sensitive component. Chrome plating may be tarnished and underlying brass corroded by chlorine bleach (sodium hypochlorite), strong acids, and concentrated oxidizers. Requires selection of neutral-pH, non-chlorine disinfectants. |

# 2.  Cleaning Procedure (Pre-Disinfection)

Cleaning must precede disinfection. Organic material (blood, body fluids, vernix) inactivates disinfectants and must be removed prior to chemical disinfection.

- PPE: Don plastic apron, disposable gloves, and eye protection. Ensure adequate ventilation.

- Debris removal: Prior to draining, remove visible debris and organic matter with a sieve or net to prevent blockage of the drain outlet. The drain plug may be removed to allow larger particles to pass.

- Drain the pool.

- Select a non-abrasive detergent: Use a pH-neutral, non-abrasive detergent approved for healthcare surfaces. Do not use scouring pads, abrasive powders, or steel wool. Abrasive products will dull the Ficore® gel coat finish and void the warranty.

- Clean tap and spray heads first: Wipe from the cleanest area (tap, spray heads, pool exterior) toward the dirtiest area (interior pool basin). This prevents transfer of microorganisms from heavily soiled surfaces to the tap and fittings.

- Rinse thoroughly: Rinse the entire pool surface with warm water. Ensure all detergent residue is removed before proceeding to disinfection.

# 3.  Disinfection Procedure

After cleaning and rinsing, proceed immediately to disinfection. Prepare disinfectant solution according to the product label at the correct use-dilution for a hospital-grade intermediate-level disinfectant claim.

- Prepare solution: Mix disinfectant per product label instructions in a clean bucket. Do not use spray bottles as the sole application method — they cannot reliably achieve 100% surface coverage. If spray application is used, follow immediately with thorough wiping to ensure full surface contact.

- Disinfect tap and spray heads first: Apply disinfectant solution to the tap and spray head assembly prior to disinfecting the pool basin, using the same clean-to-dirty directional approach.

- Wipe entire pool surface: Using a clean cloth or mop, thoroughly apply disinfectant solution across 100% of the pool interior surface, including the underside of any ledges or contoured areas.

- Maintain contact time: Allow the disinfectant to remain visibly wet on the surface for the full contact time specified on the product label for the intended pathogen kill claim. Reapply if the surface dries before contact time is met.

- Cold water rinse: Rinse the tap first, then the pool interior thoroughly with cold water to remove all disinfectant residue. This step is critical to prevent chemical residue contact with the next patient and to protect chrome-plated brass fittings from prolonged chemical exposure.

- Dry the pool: Wipe all surfaces dry with a clean, lint-free cloth.

- Close drain outlet: Keep the drain plug closed when the pool is not in use.

# 4.  Recommended Disinfectants

The following disinfectant categories are compatible with all three Active Birth Pool materials (Ficore®, stainless steel, and chrome-plated brass) when used according to label directions at recommended dilutions. All products referenced are EPA-registered for hospital/healthcare use and are widely available through US healthcare distribution channels.

RECOMMENDED The following product categories are safe for routine use on all three pool materials. ## A. Quaternary Ammonium Compounds (Quats) Neutral- to mildly-alkaline pH. Non-corrosive to metals including stainless steel and chrome-plated brass. Non-abrasive to gel-coat surfaces. Widely used as the primary surface disinfectant in US hospitals. Effective against a broad spectrum of bacteria, viruses, and fungi. Note: Always verify that the specific product label identifies stainless steel, chrome, and Ficore® among compatible surfaces. Label instructions are authoritative. ## B. Accelerated Hydrogen Peroxide (AHP) Low-level hydrogen peroxide synergized with surfactants. Breaks down to water and oxygen after use — no residual active chemical. EPA's lowest toxicity rating (Category IV). Non-corrosive to metals and Ficore® at labeled dilutions. Particularly suitable for birth pool environments given the low toxicity profile relevant to neonates and staff. Highly effective against MRSA, Norovirus, SARS-CoV-2, Candida auris, and C. difficile (product-specific; verify label). Note: Confirm product label compatibility with stainless steel and chrome-plated brass before institutional adoption, as AHP formulations vary by manufacturer.

NOT RECOMMENDED The following product types and specific products should not be used on Active Birth Pools due to risk of material damage or patient safety concerns. ## Sodium Hypochlorite (Chlorine Bleach) Bleach provides excellent broad-spectrum disinfection including C. difficile sporicidal activity, but is corrosive to both stainless steel and chrome-plated brass with repeated or prolonged use. Chloride ions disrupt the chromium oxide passive layer on stainless steel, leading to pitting corrosion. Chrome-plated brass components are particularly susceptible, as bleach can tarnish the chrome surface and corrode the underlying brass. If bleach is required by facility infection control policy (e.g., C. difficile outbreak protocols): - Use the minimum effective concentration: 1,000 ppm available chlorine (approximately 1:50 dilution of 5.25% household bleach) for standard disinfection; 5,000 ppm for C. difficile. - Apply to the pool basin (Ficore® surface) only, avoiding direct contact with the tap, spray head, and chrome-plated brass fittings to the greatest extent possible. - Do not allow bleach solution to pool or remain in contact with metal components; wipe continuously. - Strictly observe the required contact time; do not extend beyond the minimum required. - Rinse all surfaces immediately and thoroughly with cold water upon completion of contact time. Do not allow bleach to dry on metal surfaces. - Frequency should be limited to episodic outbreak management only, not routine disinfection cycles. ## Phenolic Disinfectants Not recommended for surfaces in direct contact with neonates or for immersion pools. Phenolic compounds pose a neonate safety risk — they can be absorbed through neonatal skin and may cause hyperbilirubinemia and other toxic effects. Surface residue is difficult to eliminate completely from large basin surfaces. ## Iodine-Based Products (Iodophors) Severe and permanent staining risk to Ficore® gel coat surfaces. May discolor chrome fittings. Not intended for large-surface environmental disinfection of birth pools. ## High-Level Disinfectants and Chemical Sterilants Products such as glutaraldehyde (e.g., MetriCide®), ortho-phthalaldehyde (e.g., Cidex® OPA), and high-concentration hydrogen peroxide sterilants (7.5%+) are intended for instrument sterilization, not environmental surfaces. These products are toxic if residue contacts skin or mucous membranes and are not appropriate for birth pool use. ## Acid-Based Cleaners and Descalers Strong or moderate acids (hydrochloric, citric, glycolic, oxalic acid) will corrode chrome-plated brass immediately on contact and cause pitting of stainless steel. Acid-based descalers must not be used on pool hardware. ## Abrasive Cleaners and Scouring Products Abrasive powders, cleansers, or scrubbing pads will permanently scratch and dull the Ficore® gel coat finish. Use of not recommended products voids the manufacturer warranty.

# 5.  Product Reference Tables

The following tables list example products in each category as a reference for infection control personnel. This list is not exhaustive. Always verify current EPA registration and surface compatibility on the product label prior to use.

## 5A.  Approved Products

Compatible with Ficore®, stainless steel, and chrome-plated brass when used per label directions.

| Product Name | Manufacturer | Active Ingredient | Notes |
| --- | --- | --- | --- |
| Quaternary Ammonium Compounds (Quats) — Liquids & Concentrates |  |  |  |
| Diversey Virex® II 256 | Diversey | Quaternary ammonium | Concentrate; 1:256 dilution; broad-spectrum; safe on SS, chrome, Ficore® |
| Diversey Virex® TB | Diversey | Quaternary ammonium | RTU; tuberculocidal; safe on metal and non-porous surfaces |
| Ecolab Oasis® 146 One-Step | Ecolab | Quaternary ammonium | Neutral pH concentrate; labeled for chrome, SS, plastic, Ficore® |
| Ecolab Peroxide Multi Surface Cleaner & Disinfectant | Ecolab | Quaternary ammonium / H₂O₂ | RTU; broad-spectrum; non-corrosive; hospital-grade |
| 3M™ Quat Disinfectant Cleaner Concentrate 5A | 3M | Quaternary ammonium | Neutral pH; labeled for metal, chrome, SS, and plastic surfaces |
| Spartan Chemical BNC-15® | Spartan Chemical | Quaternary ammonium | Neutral pH concentrate; hospital-grade; compatible with SS and chrome |
| Buckeye Pb 256™ Disinfectant | Buckeye International | Quaternary ammonium | Concentrate; broad-spectrum; neutral pH; safe on common hospital surfaces |
| National Chemical Laboratories Neutra-Clean D256 | National Chemical Labs | Quaternary ammonium | Low-concentration neutral quat; hospital use; safe on metals |
| Quaternary Ammonium — Ready-to-Use Wipes |  |  |  |
| PDI Sani-Cloth® AF3 Germicidal Wipes | PDI Healthcare | Quaternary ammonium (alcohol-free) | Alcohol-free; fragrance-free; 3-min contact; safe on SS, metals, plastics; EPA Reg. 9480-9 |
| PDI Super Sani-Cloth® Germicidal Wipes | PDI Healthcare | Quat + isopropyl alcohol | 1–2 min contact; broad-spectrum; compatible with SS and non-porous surfaces; EPA Reg. 9480-4 |
| PDI Sani-Cloth® Prime Germicidal Wipes | PDI Healthcare | Quat + IPA + ethanol | 1-min tuberculocidal; kills 57 organisms; compatible with SS and metal; EPA Reg. 9480-12 |
| PDI Sani-Cloth® Plus Germicidal Wipes | PDI Healthcare | Quaternary ammonium | Alcohol-free; low-level disinfectant; safe on SS and chrome; EPA Reg. 9480-3 |
| Ecolab Wipe™ N Walk Floor Disinfectant Wipes | Ecolab | Quaternary ammonium | Large-format wipes for floor and basin surfaces; hospital-grade |
| Metrex CaviWipes™ 1 | Metrex Research | Quat + isopropyl alcohol | 1-min contact; broad-spectrum; compatible with metals and plastics; widely used in US hospitals |
| Metrex CaviWipes™ XL | Metrex Research | Quat + isopropyl alcohol | Large wipe format for large surface areas; same formula as CaviWipes 1 |
| Accelerated Hydrogen Peroxide (AHP) — Liquids & Concentrates |  |  |  |
| Diversey Oxivir® 1 (RTU) | Diversey | Accelerated H₂O₂ (AHP) | 1-min virucidal/bactericidal; compatible with SS, chrome, Ficore®; EPA Reg. 70627-74 |
| Diversey Oxivir® Tb (RTU) | Diversey | Accelerated H₂O₂ (AHP) | 1-min tuberculocidal; safe on SS and chrome; EPA Reg. 70627-60 |
| Diversey Oxivir® Five 16 (concentrate) | Diversey | Accelerated H₂O₂ (AHP) | 1:16 concentrate; 5-min disinfection; EPA Reg. 70627-58 |
| Virox AHP® Rescue™ RTU | Virox Technologies | Accelerated H₂O₂ (AHP) | RTU; broad-spectrum; biodegradable; EPA lowest toxicity category; compatible with SS and chrome |
| Virox AHP® Accel™ Concentrate | Virox Technologies | Accelerated H₂O₂ (AHP) | Concentrate; hospital-grade; flexible dilutions; safe on Ficore® and metal hardware |
| Clorox Healthcare® Hydrogen Peroxide Cleaner Disinfectant Wipes | The Clorox Company | Hydrogen peroxide | 30-sec kill time; non-corrosive; labeled safe on SS, chrome, Ficore®; EPA registered |
| PDI Sani-HP1™ Germicidal Wipes | PDI Healthcare | Hydrogen peroxide | H₂O₂-based; tuberculocidal; broad-spectrum; compatible with SS and metals; EPA Reg. 9480-17 |

## 5B.  Products Not Approved for Routine Use

The following products are not suitable for use on Active Birth Pools due to material incompatibility, patient safety concerns, or other contraindications. Many are effective and appropriate hospital disinfectants for other applications but must not be used on birth pool surfaces.

| Product Name / Category | Manufacturer / Type | Active Ingredient | Reason Not Approved for Routine Use |
| --- | --- | --- | --- |
| Sodium Hypochlorite (Chlorine Bleach) — NOT for Routine Use |  |  |  |
| Clorox Healthcare® Bleach Germicidal Wipes | The Clorox Company | Sodium hypochlorite | Corrosive to chrome-plated brass and stainless steel with repeated use; causes pitting and tarnishing of metal components |
| Clorox Healthcare® Bleach Germicidal Cleaner Spray | The Clorox Company | Sodium hypochlorite | Bleach-based spray; corrosive to pool metal components; not compatible with repeated contact on SS or chrome fittings |
| PDI Sani-Cloth® Bleach Germicidal Disposable Wipes | PDI Healthcare | Sodium hypochlorite | Effective sporicidal; however bleach formulation is corrosive to SS and chrome-plated brass with routine use; EPA Reg. 9480-8 |
| Metrex CaviWipes™ Bleach | Metrex Research | Sodium hypochlorite | Bleach-based wipes; sporicidal but corrosive to chrome and SS; not approved for routine pool use |
| Clorox® Regular Bleach (household) | The Clorox Company | Sodium hypochlorite | Undiluted bleach; severely corrosive to all three pool components; not for use on medical equipment |
| Diversey Oxonia Active® (peracetic acid/H₂O₂) | Diversey | Peracetic acid + H₂O₂ | Corrosive to brass and chrome plating; may degrade gel coat with repeated use |
| Phenolic Disinfectants — NOT Approved |  |  |  |
| Lysol® Professional Disinfectant Spray (concentrate) | Reckitt | Alkyl dimethyl benzyl ammonium + phenol | Phenolic formulation; not recommended for neonatal environments; potential residue absorption; can damage gel coat with repeated use |
| Lysol® I.C. Quaternary Disinfectant Cleaner | Reckitt | Quaternary ammonium + phenolic | Contains phenolic compound; not recommended for immersion pools where neonates are present; residue risk |
| O-Syl® Disinfectant | Steris | Phenolic (o-phenylphenol) | Phenolic disinfectant; not recommended for surfaces in direct patient and neonate contact; absorbed through neonatal skin |
| Environ® One-Step Disinfectant | Steris | Phenolic compound | Phenolic chemistry; neonate safety concern in immersion birth pool environment; not approved for this use |
| Vesphene® III Disinfectant (phenolic) | Steris | Phenolic (o-phenylphenol) | Phenolic disinfectant; neonate safety concern; not suitable for birth pool surfaces |
| Iodine-Based / Iodophor Products — NOT Approved |  |  |  |
| Betadine® Solution (povidone-iodine) | Avrio Health | Povidone-iodine | Severe staining of Ficore® gel coat; may discolor chrome fittings; not intended for environmental disinfection of hard surfaces |
| Wescodyne® Plus (iodophor) | Steris | Iodophor complex | Iodine-based; staining risk on Ficore® surfaces; not recommended for immersion pool sanitation |
| General iodophor sanitizers (food-grade) | Various | Iodine complex | Designed for food contact; not appropriate for birth pool disinfection; staining and residue concerns |
| High-Level Disinfectants / Chemical Sterilants — NOT Appropriate for This Use |  |  |  |
| Metrex MetriCide® 28 (glutaraldehyde) | Metrex Research | 2.5% Glutaraldehyde | High-level disinfectant/sterilant intended for immersible instruments; toxic to staff and patients if residue remains; not appropriate for birth pool surfaces |
| Cidex® OPA Solution | Steris / J&J | Ortho-phthalaldehyde | High-level disinfectant for endoscopes; toxic if residue contacts skin or mucous membranes; not suitable for birth pool use |
| Sporox® II (hydrogen peroxide 7.5%) | Sultan Healthcare | 7.5% Hydrogen peroxide | High-concentration H₂O₂ sterilant; oxidatively corrosive to brass and chrome at this concentration; not intended for environmental surfaces |
| Acid-Based Cleaners & Descalers — NOT Approved |  |  |  |
| Lime-A-Way® (acid descaler) | Reckitt | Hydrochloric / citric acid | Acidic; corrodes chrome-plated brass immediately on contact; will pit stainless steel; not for use on pool hardware |
| CLR® Calcium Lime Rust Remover | Jelmar | Lactic / glycolic acid | Acid-based descaler; corrosive to chrome plating and brass; not for pool use |
| Tile and Tub Cleaners (HCl-based) | Various | Hydrochloric acid | Severely corrosive to all metal components; will permanently damage chrome plating and stainless steel |
| Abrasive Products — Prohibited (Voids Warranty) |  |  |  |
| Comet® Cleanser (powder) | Procter & Gamble | Sodium hypochlorite + abrasive silica | Abrasive powder; will permanently scratch and dull Ficore® gel coat finish; VOIDS manufacturer warranty |
| Ajax® Cleanser (powder) | Colgate-Palmolive | Calcium hypochlorite + abrasive | Abrasive cleanser; scratches gel coat; bleach component also corrosive to metal hardware; VOIDS warranty |
| Bar Keepers Friend® (powder) | SerVaas Labs | Oxalic acid + abrasive | Abrasive powder; acid component may etch gel coat and corrode chrome; not approved for pool surfaces |
| Scotch-Brite™ scour pads / steel wool | 3M / general | Mechanical abrasive | Mechanical abrasion; permanently scratches Ficore® surface; destroys finish and compromises surface integrity |

# 6.  Warranty Considerations & Documentation Requirements

## Non-Covered Chemical Damage

The limited warranty does not cover damage resulting from:

- Failure to rinse or properly mix disinfectants

- Allowing chemical residue to air-dry on metal components

- Use of any not approved chemicals

- Improper dwell times exceeding the label recommendations

- Use of concentrations exceeding manufacturer label instructions

- Use of non-EPA-registered (U.S.) or non-DIN-approved (Canada) disinfectants in healthcare settings

Damage examples that may be excluded from warranty coverage include:

- Pitting corrosion of stainless steel

- Chrome staining, peeling or flaking

- Surface dulling, staining, scratching, cracking, or crazing

- Chemical etching or discoloration

## Documentation Requirements

To preserve warranty coverage, facilities should:

- Maintain disinfectant product documentation (SDS and label) for products used

- Maintain cleaning protocols showing dilution and dwell times

- Document any high-level disinfection events requiring restricted or non-approved chemicals

Failure to document or prove proper chemical use as outlined in this document may impact warranty claim review.